GENERAL

sUMMARY

Section Includes

Factory Fabricated Pre-Piped and Pre-Wired Patient Care Headwall System

Patient Care Headwall as detailed on the drawings, including related components and accessories required to form an integral unit. Headwall components shown on the drawings but not specified below shall be included as part of the work under this Section, and all applicable portions of the Specification shall apply to these items.

Related Documents

Drawings and General Provisions of the Contract, including General and Supplementary Conditions and Division 1 Specification Sections, apply to this Section.

Related Sections

Section 22 60 00, Gas and Vacuum systems for Laboratory and Healthcare Facilities

Section 26 00 00, Electrical

Section 27 52 00, Healthcare Communications and Monitoring Systems

Related LEED Sections

Construction Waste Management – Section 01 7419

Sustainable Design Requirements – Section 01 8113

Indoor Air Quality Requirements – Section 01 8114

Indoor Finish Requirements – Section 01 8115

GREEN BUILDING REQUIREMENTS

The Owner requires the Contractor to implement practices and procedures to meet the project’s environmental performance goals, which include achieving the LEED Certification specified in Section Sustainable Design Requirements.  Specific project goals that may impact this area of work include: use of recycled-content materials; use of locally-manufactured materials; use of low emitting materials; construction waste recycling; and the implementation of a construction indoor air quality management plan.  The Contractor shall ensure that the requirements related to these goals are implemented to the fullest extent.  Substitutions, or other changes to the work proposed by the Contractor or their Subcontractors, shall not be allowed if such changes compromise the stated LEED Requirements.

REFERENCes

American National Standards Institute (ANSI)

American Society for Testing and Materials (ASTM)

ASTM A167-99, Stainless and Heat-Resisting Chromium-Nickel Steel Plate, Sheet and Strip

National Fire Protection Association (NFPA)

NFPA 70, National Electrical Code (NEC)

NFPA 99, Healthcare Facilities

Underwriters Laboratories (UL)

GreenGuard Certification (Components)

submittals

Submit sufficient information, clearly presented, to determine compliance with drawings and specifications, including but not limited to the following.

Product Data

Manufacturer’s Product and Utility Information Sheets.

Manufacturer’s Product Specifications.

Drawings

Elevation and Backbone/Frame drawings of Headwall showing materials, construction, dimensions, ceiling height, piping/electrical termination points, circuit designators, quantities/locations of utilities and accessories, and mounting information.

Wiring diagram showing connections, electrical ratings, and circuit designators for utilities contained within the Headwall.

As directed by Owner, rough-in locations and dimensions.

Request For Information (RFI).

Reference to available finishes and colors.

Close-out data

As directed by Owner, provide complete Operation and Maintenance Manual.

As directed by Owner, provide complete Drawings of Record.

Warranties: See Section 1.10.

quality assurance

Qualifications

Manufacturer

Has been regularly engaged in design and manufacture of the types of products and scope similar to the requirements of this project for a period of not less than five years.

Has sufficient capacity to produce and deliver required materials without causing delays in work.

Is capable of providing field service representation.

Installer

Has successfully completed at least three projects of the type and scope similar to the requirements of this project.

Can furnish proof of familiarity with equipment to be installed.

Is acceptable to Manufacturer.

Field Samples (Mock-Ups)

As directed by Owner, construct fully functional Mock-Up unit for the purpose of evaluating quality of work, operation of equipment, product application, system integration, and facility preparation.

Construct per approved drawings.

As directed by Owner, may or may not remain part of finished work.

Work may proceed only after written acceptance of Mock-Up is received.

DELIVERY, STORAGE, and HANDLING

Time deliveries to ensure materials are available on site when required.

Materials delivered to site in Manufacturer’s clearly identified containers.

Receive and store materials in their original containers in a secured space, and in a manner to prevent damage from exposure to harmful weather or conditions.

project / site conditions

Existing Conditions

Ensure that walls scheduled to receive materials are adequately reinforced to accept installation of this work.

Ensure that floor, ceiling, and wall work is completed to accept Headwall.

Report any deficiencies to Owner/Contractor for corrections.

Protection

Ensure that adjoining work will not be adversely affected by installation of work in this Section.

Provide temporary protection as required.

Environmental Limitations

Do not install materials until spaces are enclosed and weather tight, wet work in spaces is complete and dry, and temporary HVAC system is operating and maintaining ambient temperature and humidity conditions at occupancy levels during remainder of the construction period.

Field Measurements

Verify actual dimensions of construction contiguous with Headwall by field measurements before fabrication.

Field measurements to be provided by General Contractor.

Sequencing

Coordinate work in this Section with work of other trades for proper time and sequence in order to avoid delays.

Coordinate this work with other operations taking place in the same area to avoid conflicts.

WARRANTY

All warranties run from date of substantial completion.

Written warranty on entire system, signed jointly by Installer, Manufacturer, and Contractor for period of twelve (12) years.

PRODUCTS

patient bed service wall

Products

Compass™ Patient Care Modular Headwall System

Manufacturer

Hospital Systems, Inc.;  750 Garcia Avenue, Pittsburg, CA 94565 USA;  Telephone (925) 427-7800;  Fax (925) 427-0800;  Email info@HSIheadwalls.com;  Website www.HSIheadwalls.com

Single Source Responsibility: Components and materials specified in this section provided by a single Manufacturer.

If products of other Manufacturers are used to satisfy this section, all items shall meet the requirements specified herein, and be approved by Owner or Architect prior to bid.

Designed and 100% Manufactured in the USA

Design Criteria

High-quality functional Patient Care Headwall with the inherent attributes of durability, aesthetic value, and safety while providing maximum functionality.

Facilitate patient care needs in environments that require a high degree of flexibility and are subject to future changes.

Effective means to deliver necessary medical gases and other utilities to the patient.

Space efficient solution making maximum use of vertical space and floor space.

Highly organized and versatile method of storing materials unique to medical requirements.

Facilitate ease of Headwall installation and cleaning.

Regulatory Requirements

Meets or exceeds NFPA 99, 56A, 76B requirements.

Complies with NFPA 70 (NEC) requirements.

UL listed and labeled.

OSHPD approved.

Meets or exceeds local codes.

GENERAL SUMMARY

Integrated Modular Patient Care Headwall System comprised of 1 or more sections, as specified by the drawings.  Each self-supporting unit consists of a basic structural framework that hangs from Wall Rails, and removable Tiles.

materials

Horizontal Rails

Wall Rail

Anodized Aluminum.

Provides horizontal interface capability to suspend horizontal and vertical hung rail components of structural framework.

Capable of supporting static load of 250 pounds per linear foot.

Attach directly to finished 5/8” or thicker drywall.

2 Rails required, upper and lower.

Includes end caps, connectors, appropriate wall anchors.

Standard nominal widths are 24”, 36”, 48”, 60”, 72”, 84”, and 96”.

Utility Support Rail

Anodized Aluminum.

Mount between Rail-Attached Stiles to provide attachment for internal electrical/data backboxes/junction boxes and gas outlet attachment brackets.

2 Rails support 1 row of utilities.

Standard nominal widths are 24”, 36”, and 48”.

Vertical Rails

Rail-Attached Stile

18 GA Cold Rolled Steel, Black Powder Coat Finish.

Mount anywhere along Wall Rails.

2 Rails required for each section.

Supports all Tiles and components.

Creates a minimum space of 3” to accommodate plumbing, power, data, and electrical connections.

Includes Anti-Dislodgement Clips for interface to Wall Rails.

Standard nominal heights are 24” and 64”.

Chase Stile

18 GA Cold Rolled Steel, Black Powder Coat Finish.

Attach to top of Rail-Attached Stile.

2 Stiles required for attaching Chase Tile.

Includes connector kit and lockdowns.

Standard nominal height is 18”.

Tiles

Face Tile

Fire-retarding core material surfaced with a high pressure plastic laminated facing sheet.  Surface material available in Corian® and Durawrap 3D Laminate in a variety of colors.

Overlapping ship-lap design at horizontal joints to prohibit moisture and liquids from seeping (wicking) behind Tile.

Attach to Rail-Attached Stiles to provide a finished cover.

Accommodates Work Surface supports (cantilever).

Standard nominal widths are 12”, 24”, 30”, 36”, and 48”.

Standard nominal heights are 20”, 24”, and 40”.

Bed Tile

Fire-retarding core material surfaced with a high pressure plastic laminated facing sheet.  Surface material available in Corian® and Durawrap 3D Laminate in a variety of colors.

Overlapping ship-lap design at horizontal joints to prohibit moisture and liquids from seeping (wicking) behind Tile.

Attach to Rail-Attached Stiles to provide a finished cover.

Designed to take the impact of a common patient bed without additional reinforcement.

Side close-outs and soft white (LU) bottom close-out included.

Standard nominal widths are 36” and 48”.

Standard nominal height is 36”.

Utility Tile, Electrical/Data/Gas, Lower

Fire-retarding core material surfaced with a high pressure plastic laminated facing sheet.  Surface material available in Corian® and Durawrap 3D Laminate in a variety of colors.

Overlapping ship-lap design at horizontal joints to prohibit moisture and liquids from seeping (wicking) behind Tile.

Attach to Rail-Attached Stiles in the bottom position on the Stiles.

Cutouts for electrical/data/gas utilities per drawings provide easy access to backboxes and gas outlets for maintenance.

Standard nominal widths are 24”, 36”, and 48”.

Standard nominal height is 24”.

Utility Tile, Electrical/Data/Gas, Upper

Fire-retarding core material surfaced with a high pressure plastic laminated facing sheet.  Surface material available in Corian® and Durawrap 3D Laminate in a variety of colors.

Overlapping ship-lap design at horizontal joints to prohibit moisture and liquids from seeping (wicking) behind Tile.

Cutouts for electrical/data/gas utilities per drawings provide easy access to backboxes and gas outlets for maintenance.

Capable of supporting a Monitor Mounting Bracket.

Standard nominal widths are 24”, 36”, and 48”.

Standard nominal heights are 20” (Partial-Height) and 40” (Full-Height).

Chase Tile

Fire-retarding core material surfaced with a high pressure plastic laminated facing sheet.  Surface material available in Corian® and Durawrap 3D Laminate in a variety of colors.

Hides piping and conduit routed from the ceiling for gas and electrical/data utilities.

Attach to Chase Stiles.

Standard nominal widths are 24”, 36”, and 48”.

Standard 25-1/2” high Chase Tile designed for use with 64” high Rail-Attached Stiles in rooms with a ceiling height between 8’-6” and 9’-0”.

Wall-Mounted Chase

Plastic, ABS, White

Painted in the field.

Hides piping and conduit routed from the ceiling for gas and electrical/data utilities.

Attach to architectural wall.

Standard nominal width is 18”.

Standard 25” high Chase designed for use with 64” high Rail-Attached Stiles in rooms with a ceiling height between 8’-6” and 9’-0”.

components

Medical Gas Piping and Medical Gas / Vacuum Outlets

Medical Gas Piping

Type L copper pipe, cleaned, capped, and properly identified.

Factory installed and manifolded for single-point connection to building services in accordance with facility requirements and drawings.  Termination points behind Chase or above ceiling.

Factory tested.

Medical Gas / Vacuum Outlets

Location and quantity to meet the configuration of services detailed on the drawings.

Provided per Section 22 62 00 and 22 63 00.

Factory installed and tested.

Medical gas outlet Manufacturer’s standard Cover Plates.

Fixed medical gas outlets are permanently installed in one location, however, solutions to allow for future relocation or addition of outlets without requiring recertification can be incorporated during the design phase.

Electrical Wiring and Electrical Receptacles

Electrical Wiring

Normal (Standard) and Emergency (Critical) Branch Power:  #12 THHN stranded copper wire, 600 Volt, with heat resistant thermo-plastic insulation.

Ground:  #12 THHN stranded copper wire.

Factory installed for single-point connection to building power in accordance with facility requirements and drawings.  Junction boxes in top row of Utility Rails, behind Chase, or above ceiling.

Differing electrical voltages and signal wiring separated with use of EMT, Flexible Metal Conduit, or HCF cable and separated at termination points with use of dedicated junction boxes.

Factory tested.

Electrical Receptacles

Location and quantity to meet the configuration of services detailed on the drawings.

Hospital Grade 15A or 20A Decora Type.

Tamper-Resistant (Safety), Illuminated, or GFCI as directed by Owner.

Colors:  White for use on normal (standard) circuits and red for use on emergency (critical) circuits.

Factory installed and tested.

Provisions

Can accommodate provisions for, but not limited to, Nurse Call Equipment, Monitoring Equipment, Data Jacks, Phone Jacks, Lighting, etc.  Low Voltage Controller for lighting fixtures (fixture by others) can be provided and factory installed when specified.

Required EMT or Flexible Metal Conduit runs to appropriate termination point junction box factory installed.

Cover Plates provided for all provisions with the exception of Light provision and Nurse Call equipment provisions.  Trim Plates provided with Nurse Call provision.

Optional Electrical Components – Provided and Factory Installed As Shown On Drawings

Switches

Industrial Grade, 120/277 Volt, 20 Amp

Toggle, Decora, 3-Way, Momentary, etc.

Colors:  White for use on normal (standard) circuits and red for use on emergency (critical) circuits.

Grounding Jacks

Solid brass receptacles enclosed in non-conductive housing.

Spring loaded with twist-to-lock action.

Meets requirements of NEC article 517 and NFPA 99.

Secondary Circuit Breakers

Includes breaker box and door with concealed hinges for access to circuit breaker handles.

lighting (optional)

Compass Overbed Light

36” or 48” Widths

Tile-mounted at top of bed section.

Sconce Lights

Tile-mounted at top of side sections.

Trim Plate provided.

Qualux 50GX™ Exam/Reading Light

Dual-Intensity Halogen:  Exam (spot) and Reading (flood).

38-Inch knuckle-joint arm.

Qualux 50SX™ Exam/Reading Light

Dual-Intensity Halogen:  Exam (spot) and Reading (flood).

31.5-Inch gooseneck arm.

Series MLC Exam/Reading Light

Fluorescent/Incandescent.

45-Inch adjustable arm.

Series ARL Exam/Reading Light

Incandescent.

45-Inch or 50-Inch adjustable arm.

accessories (OPTIONAL)

Horizontal Eclipse™ Equipment Rails

Anodized Aluminum, Satin Oxide Laminate Insert, Grey End Caps.

Supports rail-mount accessories with the use of Universal Rail Clamps.

Attach to front of Tile.

Standard lengths are 18”, 24”, 30”, 36”, 42”, and 48”.

Large selection of rail-mount accessories available including Vacuum Bottles and Slides, Monitor Mounting Channels and Arms, Diagnostic Equipment, Sharps Containers, Glove Holders, I.V. Poles, Soap Dispensers, Chart Holders, Utility Hooks, etc.

Accessories to be provided by Manufacturer of Headwall to ensure compatibility.

EXECUTION

examination

Ensure structure and/or substrate is adequate to support Headwall.

Inspect and verify that walls and areas in which work is to be performed are acceptable for Headwall installation in accordance with Manufacturer’s published recommendations and all applicable Sections.

Report all discrepancies and unacceptable conditions to Owner/Contractor for corrections.

Proceed with installation only after discrepancies and unacceptable conditions have been remedied.

preparation

Coordinate Headwall installation with work of other trades for proper sequence to avoid delays.

Coordinate service connection work with Electrical, Piping, and Communication Contractors.

INSTALLATION

Perform installation in strict accordance with Manufacturer’s published recommendations and all applicable Sections.

Anchor all fixed components securely, square, level, and plumb at heights indicated on drawings.

Align slots in vertical support elements to ensure hanging units are level.

After installation of Headwall and connection of services are complete, test electrical equipment function and grounding in accordance with NFPA requirements.  Purge and test medical gases for system certification in accordance with NFPA 99.

cleaning

Clean all surfaces to remove all marks, soil, and foreign matter immediately after installation and adjustment are complete.

Recheck all components and perform any necessary additional cleaning just prior to substantial completion.

Remove surplus materials, debris, tools, and equipment upon completion.

Manage Packaging Waste

Collect and separate packaging waste materials for reuse and/or recycling.

Remove packaging materials from site and dispose of at appropriate recycling facilities.

protection

Protect installed Headwall from damage during remaining construction work.

END OF SECTION