EMI and EMC Pre-Compliance Set up and Training kit This System should have the Capability for EMC Pre- compliance Testing The Training and Pre-Compliance System should consist of the following 1. Signal Analyzer with software for EMI test report management and assessment and facility to select the required CISPR EMI regulation for conduction and radiation measurement. 2. EMI Near Field Probe Kit 3. Two Line V Network with Transient Limiter 4. Isolation Transformer 5. RF Circuit Training Module 6. Microstrip Component Training Module Signal Analyser: Frequency Range: 9kHz ~ 3.25GHz with 0.025ppm Frequency Stability and 1ppm aging Rate RBW: 1Hz ~ 1MHz (3dB), 6dB EMI Filter: 200Hz, 9kHz, 120kHz, 1MHz Fastest sweep time: 204us Sensitivity: -149 dBm/Hz (@PreAmp on) Built-in Preamplifier, 50dB Attenuator, and Sequence Function Built-in EMC pretest function, Quasi-Peak/Average EMI detect mode Built-in 2FSK Analysis, AM/FM/ASK/FSK Demodulation & Analysis Built-in P1dB point, Harmonic, CHPW, OCBW, ACPR, SEM, TOI, CNR, CTB, CSO, N-dB bandwidth, Noise Marker, Frequency Counter, TDP, Gated Sweep Built-in Spectrogram, Topographic and Split-Window Display Modes Remote control EMI measurement software Remote Control Interface: LAN, USB, RS-232 EMI Near Field Probe Kit EMI Near Field Probe Set should have four probes. The antenna factors of these four probes should be built in the EMC Pretest function of the analyzer. Two Sensing Probes are EMI magnetic field probes, which can maximally sustain 50Vdc.Two Contact Probes of which one is an AC voltage probe, which can maximally sustain CAT II, 300VAC. And Other is an EMI electric field probe, which can maximally sustain CAT I 50V DC. Two Line V – Network with Transient Limiter Two-Line V-Network(LISN) should meet CISPR16-1-2:2006 standard which is required for EMI conduction testing. Isolation Transformer RF Circuit Training Module Consisting of Transmission channels and microstrip discontinuities • Open air transmission • Coaxial cable transmission • Microstrip transmission line • CPW transmission line (with and without ground) • Slot line • Coupled line microstrip transmission line • Various discontinuities (open, short, gap, bend, step, tee) RF Circuits and Matching Architecture • RF lump elements (capacitive, inductive, matched load conditions) • High-frequency effects on lump elements • Resonance circuits (inductive, capacitive, RLC, tunable) • BALUN • Bias tees (DC, microstrip) • Matching networks (single stub, binomial quarter transformer, L, T, PI) RF Active Components Key RF active components should be available, such as: • RF BJT amplifier • Low noise amplifier • RF power amplifier • RF MOSFET amplifier • RF gain equalizer • RF LC oscillator • Phase-locked loop (PLL) • RF frequency multiplier • Lumped filters (LPF, HPF, BPF) Microstrip Component Training Module This module should feature a board with both passive and active components, including: • Filters (LPF, HPF, BPF, BSF) • Couplers (branch line, coupled line, rat race) • Attenuators (Tee, Pi) • Circulator-isolator • Power divider • Ring resonator • Transmission lines with terminations and loads • Active circuits (amplifier, mixer, VCO, phase shifters, detectors, modulators, switches, variable attenuators)